| Experiment Proposal: AR_2013_No1 |  |
| --- | --- |
| Title | Author/Spokesperson |
| Test of Ionisation Chambers & QFW electronics with Fast & Slow Extraction | A. Reiter (1431) M. Witthaus (2447) |
| Summary & Aim |  |
| - R&D for FAIR: Adaptation of QFW electronics to fast & slow extraction - Feasibility of air-IC application in high-radiation environment at pBar Separator - The response of an ionisation chambers will be investigated for short beam pulses. These pulses will be produced from SIS18 using fast extraction (h=4 or h=1). - Requested beam time: 5 shifts - Measured quantities: Signal strength and linearity (comparison/extrapolation to pBar Separator parameters) - Test of DAQ & analogue electronics chain: QFW & pulse stretcher - Recorded data: Chamber signals as function of particle number, QFW/pulse stretcher parameters and high voltage setting |  |
| Machine parameters |  |
| Machine; SIS18, fast extraction, h=4 or h=1. Mode; B-exp. Exp. area; HTP. Ion species; Uranium, Nitrogen. Beam energy; 300 – 800 MeV/u, in principle any. Spill length; Slow extraction: 300 ms – 1000 ms Fast extraction: 1 µs extraction; 4 bunches of ~100 ns length each or single bunch (h=1). Particle number; 104 – 109 per spill. Repetition rate; ~0.5 Hz or higher. Shifts; 5 shifts. Beam Time Period; Any machine experiment (B-exp) after June 2013 Several measurement blocks. Health & Safety; No concerns. |
| Experiment procedure |  |
| After setup and optimisation of beam line to HTP, the following series of data are taken: - Ionisation chamber responses as function of particle number, ionisation production rate, high voltage setting - QFW: minimum detectable signal, linearity - Reference detectors at HTP: Ionisation chamber HTPDI1I and SEM HTPDI1S or transformers RT and FCT - DAQ test of readout electronics: Charge-to-frequency converter (QFW) and scaler/counter system |  |

Experiment Setup

Exp. area; HTP, beam dump area. Description of setup; - SEM-type ionisation chamber (IC) in air + several LHC-type BLMI + HTPDI1I + FAIR-prototype MWPC/IC - Detector positions behind exit of last diagnostics chamber or in beam line (FAIR prototype chamber + HTPDK1) - Distance to beam dump: ~ 50-100 cm - Test IC with positive and negative HV. Duration of setup; - 2013/2014 measurement campaign. DAQ & Electronics Software; DAQ setup in Atomic Physics container, signal stretcher (RC-circuit), QFW or IFC converter & Ablass/Lassie FESA system, digital oscilloscope Standard operating software + Java GUI for FESA DAQ. Trigger; Slow extraction: Cycle start – cycle end (acquire full spill) Fast extraction: Flattop reached.

Experiment Preparation / Required support Estimated amount of time, manpower and equipment

Estimates or simulations; 1 day; Signal estimate A. Reiter. Mech. Workshop; Not required. Beam Line Installation; 3 h; H. Graf: Remount SEM on pneumatic actuator, alignment. Electronics & DAQ; 3 days; Setup of electronics in DAQ container and tests (A. Reiter / M. Witthaus). Control System Integration; 1 day; Trigger of Ablass scaler system: Due to pulse stretching acquisition between events [32,55] ??? Window: [flat top, extr.+250 ms]. On-site tests. Modification of exp. area; No; Pneumatic actuator in place at HTP 2 RG214 cables to AP container from patch panel. Dismantling of setup; 2 h; Dismount setup, store detector at HTP (A. Reiter).

Remarks & Comments