| Experiment Proposal: AR_2013_No11 |
| --- |
| Title | Author/Spokesperson |
| Test of Position-Sensitive BPM-type Ionisation Chamber & QFW electronics with Fast & Slow Extraction | A. Reiter (1431) M. Witthaus (2447) |
| Summary & Aim |
| R&D for FAIR: Feasibility of position-sensitive Air-IC application in high-radiation environment at pBar Separator. Monitoring of beam position in front of pBar Target Station. The response of an ionisation chambers will be investigated for short beam pulses. These pulses will be produced from SIS18 using fast or slow extraction (h=4 or h=1). Requested beam time: 1 shifts other tests in parasitic mode Measured quantities: Signal strength and linearity (comparison/extrapolation to pBar Separator parameters), position dependence of signal difference Test of DAQ & analogue electronics chain: QFW & pulse stretcher Recorded data: Chamber signals as function of particle number, QFW/pulse stretcher parameters and high voltage setting |
| Machine parameters |
| Machine SIS18, fast extraction, h=4 or h=1 Mode B-exp Exp. area HTP Ion species Uranium, Nitrogen Beam energy 300 – 800 MeV/u, in principle any Spill length Slow extraction: 300 ms – 1000 ms Fast extraction: 1 µs extraction; 4 bunches of ~100 ns length each or single bunch (h=1) Particle number 104 – 109 per spill Repetition rate ~0.5 Hz or higher Shifts 1 shifts Beam Time Period Any machine experiment (B-exp) after June 2013 Several measurement blocks Health & Safety No concerns |
| Experiment procedure |
| After setup and optimisation of beam line to HTP, the following series of data are taken: Ionisation chamber responses as function of particle number, ionisation production rate, high voltage setting Position dependence of signal difference can be evaluated by changing the vertical detector position with a stepper motor. QFW: minimum detectable signal, linearity Reference detectors at HTP: Ionisation chamber HTPDI1I and SEM HTPDI1S or transformers RT and FCT DAQ test of readout electronics: Charge-to-frequency converter (QFW) and scaler/counter system |

Experiment Setup

Exp. area HTP, beam dump area Description of setup Segmented BPM-type ionisation chamber (IC) in air + HTPDI1I/HTPDI1S Detector positions behind exit of last diagnostics chamber or in beam line (FAIR prototype chamber or HTPDK1) Distance to beam dump: ~ 50-100 cm Test IC with positive and negative HV Duration of setup 2013/2014 measurement campaign DAQ & Electronics Software DAQ setup in Atomic Physics container, signal stretcher (RC-circuit), QFW or IFC converter & Ablass/Lassie FESA system, digital oscilloscope Standard operating software + Java GUI for FESA DAQ Trigger Slow extraction: Cycle start – cycle end (acquire full spill) Fast extraction: Flattop reached

Experiment Preparation / Required support Estimated amount of time, manpower and equipment

Estimates or simulations 1 day Signal estimate A. Reiter, like IC Mech. Workshop Not required Beam Line Installation 3 h H. Graf: Mount detector on vertical stepper actuator, rough alignment wrt. beam axis Electronics & DAQ 3 days Setup of electronics in DAQ container and tests (A. Reiter / M. Witthaus) Control System Integration 1 day Trigger of Ablass scaler system: Due to pulse stretching acquisition between events [32,55] ??? Window: [flat top, extr.+250 ms] On-site tests Modification of exp. area No Vertical stepper-motor actuator at HTP 2 RG214 cables to AP container from patch panel Dismantling of setup 2 h Dismount setup, store detector at HTP (A. Reiter)

Remarks & Comments