## Slide 1
M. Almalki
(Beam Diagnostic Group GSI, Goethe-Universität)
Deutsche Physikalische Gesellschaft ( 04 – 08 March 2013 )
Digital Beam Position and Phase Monitor for p-LINAC for FAIR
M. Almalki1, P. Forck1, W. Kaufmann1, P. Kowina1, R.Singh1, C. Krüger1,
B. Baricevic2, M. Znidarcic2, R. Hrovatin2
1GSI, Darmstadt, Germany; 2I-Tech, Solkan, Slovenia
1-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR

## Slide 2
Purpose
The results of the beam-based test for digital beam position determination and beam phase detection using Libera SPH (Libera Single Pass Hadron).
Libera SPH
2-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR

## Slide 3
Motivation (Beam Position and Beam Phase).
1
Outline
Beam-based Test with Libera SPH.
3
Digital Signal Processing Algorithm for Position and Phase.
2
Results & Conclusion.
4
3-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR

## Slide 4
Outline
4-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Beam-based Test with Libera SPH.
3
Digital Signal Processing Algorithm for Position and Phase.
2
Results & Conclusion.
4
Motivation (Beam Position and Beam Phase).
1

## Slide 5
Circumference
1080 m
7 × 10 cooled pbar/hour
10
2 × 10 proton/hour
16
5-33
Motivation
1
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR

## Slide 6
Circumference
1080 m
2 × 10 proton/hour
16
Motivation.
1
6-33
7 × 10 cooled pbar/hour
10
48 60 70 MeV
3 MeV 24 36
CH4 to CH6
CCH1 to CCH3
Source LEPT RFQ
95 keV
dump
to SIS18
30 m
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
p-LINAC for FAIR

## Slide 7
Beam Position Beam Phase
BPM 1
BPM 2
L
Accelerator
7-33
Beam Position and Beam Phase Detection
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
∆t
Position
Position
Phase
Beam energy from time-of -flight (TOF):
accuracy of 1 = 8.5 ps @ 325 MHz
o
Position :
0.1 mm spacial resolution.

## Slide 8
8-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Beam-based Test with Libera SPH.
3
Results & Conclusion.
4
Motivation (Beam Position and Beam Phase).
1
Digital Signal Processing Algorithm for Position and Phase.
2

## Slide 9
RF analogue front-end
Sampled at
117 MHz
IF (29.4 MHz)
Signals from BPM
Libera SPH
Digital Signal Processing Algorithm in p-LINAC
2
9-33
Digital processing (IQ demodulation)
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
325.2 MHz
- I
I
- Q
Q

## Slide 10
Extracting I and Q (amplitude and phase )
NCO
LPF
LPF
I
Q
Local Oscillator
Digital Signal Processing Algorithm in p-LINAC
2
10-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
I : in-phase component
Q : quadrature-phase component
RF application in DSP,
T. Tschilcher, CAS 2007
Filter out
Filter out
Amplitude
Phase

## Slide 11
11-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Results & Conclusion.
4
Motivation (Beam Position and Beam Phase).
1
Digital Signal Processing Algorithm for Position and Phase.
2
Beam-based Test with Libera SPH.
3

## Slide 12
1- Beam-based Test (Purpose and Method).
2- Experimental Setup (Beam Line Schematic).
3- Data Treatment for different bunch shapes.
4- Case Study (as an example)
Beam-based Test for Digital Beam Position and Phase Monitor Using Libera SPH
12-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR

## Slide 13
Beam-Based Test (Purpose and Method).
1
13-33
Observation:
Different bunch shape leads to different signal form (shape)
different signal form different arriving time.
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR

## Slide 14
Beam-Based Test (Purpose and Method).
1
14-33
The Aims are :
1- to test beam position accuracy.
2- to characterize the dependence of beam arrival time on bunch shape
Libera SPH
Fast Oscilloscope (20 GSa/s )
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Observation:
Different bunch shape leads to different signal form (shape)
different signal form Different arriving time.

## Slide 15
BPM
Accelerator
Position
Phase
(Zero Crossing)
Method :
Generating different bunch shapes and a single BPM acts as a "Bunch arrival monitor".
BPM
Beam-Based Test (Purpose and Method).
1
15-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Observation:
Different bunch shape leads to different signal form (shape)
different signal form Different arriving time.
Libera SPH
Oscilloscope & Libera SPH

## Slide 16
1
2
BPM
Beam
1 - Buncher for bunch gymnastics longitudinally
2- Quadrupole doublet for beam focussing
3- Dipole for position variation horizontally
BPM : Capacitive pickup with four electrodes.
44 dB amplification
Experimental Setup (Beam Line Schematic).
2
16-33
3
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Bunch rep. rate | 108.408 MHz
Macropulse | 200 μs
Macropulse rep.rate | 1 Hz
Beam Current | 60 ~ 90 μA
Beam Energy | Ca10+ @ 1.4MeV /u
The Macropulse structure

## Slide 17
Experimental Setup (Beam Line Schematic).
2
17-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
BPM
Bunch rep. rate | 108.408 MHz
Macropulse | 200 μs
Macropulse rep.rate | 1 Hz
Beam Current | 60 ~ 90 μA
Beam Energy | Ca10+ @ 1.4MeV /u
The Macropulse structure

## Slide 18
Seven Bunch Shapes
3
18-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR

## Slide 19
For each bunch shape
The amplitude was measured in steps of 1 μs (correspond to 108 bunches) for a train of 129 μs.
S = S = 14 % / mm
Beam Position
Horizontal axis
Vertical axis
y
x
Data Treatment - From Libera SPH -
3
20-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
One macropulse

## Slide 20
The " Libera SPH" provides the measured phase in respect to a master oscillator – RF accelerating frequency.
Signal Phase
Data Treatment - From Libera SPH -
3
21-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
For each bunch shape
One macropulse

## Slide 21
Amplitude (V)
2 ns/div
500 Mv/div
Data Treatment - In time-domain (20 GSa/s Oscilloscope)
3
22-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
For each bunch shape
Zero Crossing
The signal is interpolated (spline interpolation)
213 points

## Slide 22
Data from 20 GSa/s Oscilloscope, for each bunch shape
Data Tratment - FFT of a single Bunch -
3
23-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Extracting window for only one bunch for FFT.
-116.8ͦ
210 points
215
215 points
FFT calculation of a single bunch.

## Slide 23
Data from 20 GSa/s Oscilloscope, for each bunch shape
Data Tratment - FFT of a train of bunches -
3
24-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
A train of bunches (25000000 samples)
1st harmonic 2nd harmonic 3rd harmonic
30000 bunches

## Slide 24
To remaind,
To compare phases using two different techniques.
Libera SPH (digital processing) &
Oscilloscope (time-domain)
Time-domain data are Fourier-transformed to calculate amplitude and phase
25-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
position & phase
phase
Spectrum of amplitude & phase

## Slide 25
-10.22 ͦ (0.28)
-17.74 ͦ (0.72)
147 ps
27 ps = 1 ͦ @ 108.4 MHz
5.74 ͦ
7.86 ͦ
Shape 3
Reference shape
26-33
Case study (as an example )
4
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Time (ns)
-50 0 50 100 150 200 300
Time (ps)

## Slide 26
-110.07 ͦ
-116.85 ͦ
-109.24 ͦ
-117.08 ͦ
6.7ͦ
7.8ͦ
27-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Shape 3
Reference shape
Time (ns)
4
Case study (as an example )

## Slide 27
Libera SPH
Oscilloscope
A single
bunch
FFT calculations
A train
of bunches
5.74 ͦ
7.86 ͦ
7.8 ͦ
6.7 ͦ
28-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Shape 3
Reference shape
Time (ns)
4
Case study (as an example )

## Slide 28
Beam-based Test with Libera SPH.
3
29-33
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Motivation (Beam Position and Beam Phase).
1
Digital Signal Processing Algorithm for Position and Phase.
2
Results & Conclusion.
4

## Slide 29
30-33
Results
5
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Phase (degree)
Zero Crossing t (degree @ 108 MHz)
Beam Phase Detection

## Slide 30
Results
5
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Beam Position Determination
X : Horizontal plane
Y : Vertical plane
31-33

## Slide 31
Conclusion
5
Beam-based Test (purpose, method, setup).
Libera SPH phase evaluation in comparison with time-domain and FFT calculations.
1- The tested Libera SPH digital electronics is siuted for beam position determination.
2- For phase detection it does not fulfil the requirements yet. However, modifications might lead to improve the performance in order to meet with the desired accuracy.
What have been done ?
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
Outcomes :
32-33

## Slide 32
Thank You
M. Almalki , Dresden 2013 - DPG Digital Beam Position and Phase Monitor for P-LINAC for FAIR
33-33