Technical Concept

| Component: | Scintillating Screens | PSP: | 2.8.6.6.1.6 SIS100 2.3.6.5.2.6 HEBT 2.5.6.6.1.6 CR 2.9.6.3.2.6 pBar 2.11.6.8.6 HESR |
| --- | --- | --- | --- |
| Document: | F-TC-BD-DAQ_SCR |  | Version 1.0 28.05.2019 |
| Authors: | T. Hoffmann, B. Walasek-Höhne, C. Schmidt, A. Petit, H. Bräuning, C. Dorn |  |  |

Table of contents

1	Scope	2

2	Introduction and Instrument Purpose	2

3	Schematic System Overview and Locations	4

4	Instrument Hardware	5

4.1	Detector Hardware	5

4.1.1	Pneumatic Driven Linear Actuator	5

4.2	Optical Setup	6

4.2.1	Screen material	6

4.2.2	Viewports of sufficient optical quality	7

4.2.3	Lens system with remote-controlled iris	7

4.2.4	Camera system	8

4.2.5	Optical shielding for the camera-system to avoid incidence of ambient light  and LED unit for target illumination during calibration	9

4.3	Additional Parts / Assembly	10

4.4	Components at the beam line:	11

4.4.1	Components in the tunnel:	11

4.4.2	Components in the electronics room:	11

4.4.3	Cabling:	11

4.5	Data Acquisition Hardware	12

4.5.1	CMOS and radiation-hardened cameras	12

4.5.2	µTCA	13

4.5.3	Industry PC (IPC)	13

4.5.4	Remote Control	14

4.5.5	CPS8	14

4.5.6	PLC System	14

4.6	Supported Infrastructure (HV, Pressurized Air, Detector Gas)	16

5	Interfaces	16

5.1	Interface to Accelerator Control System	16

5.2	White Rabbit Timing	16

5.3	Beam Transmission Monitoring	16

5.4	Machine Protection	16

5.5	Accelerator and Beam Modes	16

5.6	Archiving	17

5.7	Alarming	17

5.8	Expected Data Supply from LSA	17

5.9	Miscellaneous (Interlock etc.)	17

6	Instrument Front-End Software	17

6.1	High-Level Design	17

6.2	Device Drivers and Firmware	17

6.3	Operating System	17

7	Expected Performance	17

7.1	Important Performance Parameters	17

7.2	Expert-GUI	18

8	Bill of Material	19

8.1	Following CUPID DAQ components are required:	19

8.1.1	Following CUPID µTCA DAQ components are required:	19

8.1.2	Following CUPID IPC DAQ components are required:	19

8.2	Following PLC Slow Controls components are required for each electronic room	20

8.3	Following CUPID CPS8 components are required for each electronic room	21

8.4	Following CUPID IDS digital cameras are required for each electronic room	22

8.5	Following CUPID MRad3 Cameras and Pleora frame grabber are required for each electronic room:	23

9	References	23

10	Open Issues	24

11	List of Abbreviations	25

# Scope

This technical concept shall provide information about the envisaged setup of the data acquisition hardware in combination with the detector and all required accessories such as pre-amplifiers, converters and adapter boxes of the described beam instrumentation device. In addition, it shall provide information about integration details such as location, trigger and timing issues, dependencies to other systems and open issues.

# Introduction and Instrument Purpose

At FAIR and GSI several scintillating screens are used. Those types of detectors are intercepting diagnostic devices for the determination of the beam profile.

The relevant specification can be accessed here: F-DS-BD-16e_OptProfile_SCR_T [1].

# Schematic System Overview and Locations

Figure 1: Scintillating screen system schematic.

Table 1: FAIR List

| Accelerator | Form-Factor | Device Name | Beam line / Location | Electronic Room / Rack No. |
| --- | --- | --- | --- | --- |
| HEBT |  |  |  |  |
|  | µTCA | T1S1DF2 T1S1DF3 T1S1DF5 T1S1DF6 | T1S1 T1S1 T1S1 T1S1 | G004a G0704A.E20.010 |
|  | µTCA | T1S4DF2 T1X1DF1(Radhard) T1C1DF1 T1D1DF1 TSX1DF2 TSX2DF1 TXL1DF1 TXL2DF2 SIS18 Dump (Radhard) SIS100 Dump (Radhard) | T1S4 T1X1 T1C1 T1D1 TSX1 TSX2 TXL1 TXL2 T8DU T1D1 | G017.1 H0719A.E30.123 |
|  | µTCA | TXL2DF4 TSN1DF1 TSN1DF3 | TXL2 TSN1 TSN1 | G018 L0516A.E10.130 |
|  | µTCA | THS1DF2 THS1DF4 THS1DF5 TFC1DF1 | THS1 THS1 THS1 TFC1 | G007 H0209A.E10.044 |
| SIS100 |  |  |  |  |
|  | IPC | 1S61DF1 | Section 6 | K923A.U30.160 |
|  | IPC | 1S53MP1L | Section 5 | K923A.U30.134 |
| CR |  |  |  |  |
|  | µTCA | CR01DF1S CR02DF1S CR02DF2S CR04DF1S CR04DF2S | Im Zuständigkeit von Budker Institut | G007 H0209A.E10.044 |
| CR Injection |  | FREXDF2 TFC1DF TCR1DF TCR1DF TCR1DF | FREX TFC1 TCR1 TCR1 TCR1 | G007 H0209A.E10.044 |
| pBar |  | PT01DF1 (Radhard) PS01DF7 PS01DF8 | PT01 (Target) PS01 (Separator) PS01 (Separator) | G006c K0321A.E30.016 G007 H0209A.E10.044 G007 H0209A.E10.044 |
| HESR |  | HR01DF1 HR07DF1 HR17DF1 THS1DF6 HR14DF1 | Im Zuständigkeit von FZJ Institut | HESR Süd HESR Nord HESR Süd HESR Süd HESR Nord |

# Instrument Hardware

## Detector Hardware

The scintillating screen is inserted into the beam under an angle of 45°. Through a viewport, located at 90° with respect to the beam, a camera observes the screen.

Figure 2: The drawing of the scintillating screen setup on a flange with a target under 45° in respect to the beam.

The mechanical constituents of one Scintillating Screen Setup are the following parts:

- Detector consisting of scintillating screen and target holder.

- Pneumatic driven linear actuator.

- Diagnostic chamber for installation of linear actuators.

- Additional parts for detector installation on linear actuator.

### Pneumatic Driven Linear Actuator

The linear actuator moves the detector (scintillating screen and holder) into the beam center and out of the beam. The four different detectors are foreseen correspond to four different linear actuators

| Detector | Distance Flange- Beam center [mm] | Stroke [mm] | Bake-out | BD-type number | Drawing number |
| --- | --- | --- | --- | --- | --- |
| FT-DF 2010 | 270 | 150 | No | FT-DL 2210 | BD-1018821-A-000 |
| FT-DF 2011 | 270 | 150 | Yes | FT-DL 2220 | BD-1018881-A-000 |
| FT-DF 2020 | 370 | 200 | No | FT-DL 2230 | BD-1014852-A-000 |
| FT-DF 2021 | 370 | 200 | Yes | FT-DL 2240 | BD-1018450-A-000 |

Figure 3: Technical drawing of linear actuator with camera setup.

## Optical Setup

The optical constituents of one Scintillating Screen are the following parts:

### Screen material

The detector consists of the scintillating screen (target) and the target holder. The scintillating screen is a round disk of scintillating material (solid or coated material), fixed in a holder with permanent alignment marks outside the active area.

Figure 4: The technical drawing (left) and photo (right) of the scintillating screen on target holder.

| SCR Material | Type | Application |
| --- | --- | --- |
| Al2O3:Cr (Chromox) | Solid target material | Standard light applications |
| Gd2O2S:Tb (P43) | Coated on a plate | Low light applications |
| YAG:Ce | Solid target material | Low light applications |

In the HEBT lines, two different detector sizes are used with an active area of Ø 100 mm and Ø 120 mm, mainly Chromox.

| Detector size (active area) | Bake-out | BD-type number | Drawing number |
| --- | --- | --- | --- |
| D100 | No | FT-DF 2010 | BD-1018552-A-000 |
| D100 | Yes | FT-DF 2011 | BD-1018926-A-000 |
| D120 | No | FT-DF 2020 | BD-1017884-A-000 |
| D120 | Yes | FT-DF 2021 | BD-1018471-A-000 |

### Viewports of sufficient optical quality

For standard location CF viewport with firmly connected borosilicate window VPCF16B-L (for LED Unit) and VPCF40B-L (Camera system) from VACOM will be used.  Flange from SS 1.4307 (304L) and spinning material: Kovar are used. Please keep in mind that that only annealed copper gaskets can be use.

Figure 5: Photo of viewports for scintillating screen system.

For bakeout location CF viewport with connected borosilicate window Zero Light Glass HOVPZ16-LN for LED and HOVPZ38-LN for camera system from company Hositrad can be used. Only flange from material 316L can be use. Please keep in mind that that only annealed copper gaskets can be use.

### Lens system with remote-controlled iris

The standard lens installed is the PENTAX C1614ER with fixed focal length (16mm) or PENTAX C814ER with fixed focal length (8mm), low distortion and good optical quality. A lens system with remote-controlled iris is required in order to adjust the amount of light that impinges onto the digital camera sensor. Note that the iris setting affects the depth of field.

The remote control is done by setting a defined analogue input voltage as a set value for the iris opening.

Figure 6: Photo of remote-controlled iris Pentax C1614ER.

Pentax -ER Series

- iris control by DC voltage (1.5 – 5.5V)

- iris closes when power is cut off

- manual focus

### Camera system

The standard camera installed is the IDS uEye UI-5240SE-M-GL, which is a digital GigE camera equipped with the radiation tested e2v CMOS sensor with 1280 by 1024 pixels. A camera internal area-of-interest limits the image to the size of the scintillating screen and reduces the amount of data transferred already in the camera. Grouped according to their location, the cameras are connected via 1 GbE Ethernet switch to DAQ System. The DAQ System provides a private network for the camera readout. The DAQ System also runs the FESA classes and provides the link via the accelerator network to the GUI.

The performance of the system reached more than 15 frames per second at full resolution with one active camera and a single connected client.

Figure 7: Photo of CMOS Camera IDS uEye UI-5240SE-M-GL.

For  installations, where a high radiation level is observed, a radiation-hardened solid-state CID (Charge Injection Device) based camera from Thermo Fischer Scientific, the CCIR MegaRad3 (8726DX7), will be installed. According to the manufacturer this device is tolerant to gammas, neutrons, high energy electrons and proton radiation to at least 3 MRad. Only in tests up to 14 MRad a noticeable degradation in the image quality has been reported by the manufacturer. Tests at GSI SIS18 extraction with the MegaRad3 camera have shown, that within few years of operation, no significant change in camera performance was observed.

Figure 8: Photo of radiation-hardened CID Camera MegaRad3 (8726DX7).

Dedicated FESA class was development to acquire images from a frame grabber card Pleora (Model PT01-AN1IP01-32EG) to be used with the radiation hardened camera.

Grouped according to their location, the radiation hardened cameras from Thermo Fischer Scientific are connected via frame grabber card Pleora (Model PT01-AN1IP01-32EG) to the GbE Ethernet switch inserted in the DAQ System. The DAQ System provides a private network for the camera readout.

Figure 9: Photo of frame grabber card Pleora (Model PT01-AN1IP01-32EG).

Pleora iPort Analog-Pro External Frame Grabber

- Gbit ethernet based

- GigE Vision and GenICam compatible

Analog cameras run continuously

- rate reduction in software to 10Hz

- trigger by software

- all frames are discarded but next frame after trigger

### Optical shielding for the camera-system to avoid incidence of ambient light  and LED unit for target illumination during calibration

The following drawings show the parts and material to assemble the camera shielding and LED unit on the linear actuator.

Figure 10: The technical drawing of camera shielding and LED unit assemble.

|  | Drawing number |
| --- | --- |
| LED Unit | BD-1015625-A-000 |
| Camera Unit | BD-1015564-A-000 |

| BD-type number | Detector | Linear actuator | Drawing number |
| --- | --- | --- | --- |
| FT-DZ 2211 | FT-DF 2010 | FT-DL 2210 | BD-1018825-A-000 |
| FT-DZ 2221 | FT-DF 2011 | FT-DL 2220 | BD-1018966-A-000 |
| FT-DZ 2231 | FT-DF 2020 | FT-DL 2230 | BD-1015663-A-000 |
| FT-DZ 2241 | FT-DF 2021 | FT-DL 2240 | BD-1018484-A-000 |

Figure 11: Photo of linear actuator with mounted camera, lens and optical shielding.

## Additional Parts / Assembly

Additional parts for the beam diagnostic assembly is connector box with fixing options.

Figure 12: Photo of adapter box.

Figure 13: Photo of complete mechanical and optical system for scintillating screen.

## Components at the beam line:

Besides the detector mechanics and scintillating screen, a camera with lens system and connector box is installed at the beam line. Long cables (copper) lead to the DAQ and PLC systems.

### Components in the tunnel:

If unavoidable, the PLC system for control of the lens system and LED can be mounted in a 19” rack system in the tunnel in radiation protected/reduced areas.

### Components in the electronics room:

All components of the SCR DAQ system as well as the required PLC infrastructure should be installed in electronics rooms. They are mounted in 19” rack systems.

### Cabling:

#### GbE Version

Long cables:

1x halogen free UNITRONIC S/FTP CAT7 LSZH 4x2xAWG23 network cable for camera data signals

1x halogen free DATAFLAMM-C 5x0.14mm2 between PLC and Adapter box

1x halogen free DATAFLAMM-C PAAR 5x2x0.14mm2 between CPS8 and Adapter box

Short cables:

1x halogen free DATAFLAMM-C 5x0.14mm2 between Adapter box and lens

1x halogen free DATAFLAMM-C 2x0.14mm2 between Adapter box and LED

#### Radiation-hardened Version

Long cables:

1x remote head cable, LS0H for camera data signals

1x halogen free DATAFLAMM-C 5x0.14mm2 between PLC and Adapter box

Short cables:

1x halogen free DATAFLAMM-C 5x0.14mm2 between Adapter box and lens

1x halogen free DATAFLAMM-C 2x0.14mm2 between Adapter box and LED

## Data Acquisition Hardware

At FAIR the decision for a DAQ system is based on availability on the market, costs, synergy effects, and performance.

Two possible solutions for SCR DAQ are envisaged. For installations with up to two GbE camera systems the IPC as DAQ host may be used. With higher amount of channels a µTCA solution, 6 slots, is proposed. Those systems will contain an 8 channel GbE switch.

Figure 14: Typical CUPID DAQ installation.

### CMOS and radiation-hardened cameras

| CMOS Camera requirements |  |  |
| --- | --- | --- |
| Type |  | IDS uEye UI-5240SE-M-GL CMOS digital camera, monochrome |
| Data interface |  | Gigabit Ethernet |
| Chip size | inch | 1/8’’ |
| Chip resolution | pixel | 1280x1024 |
| Refresh rate | fps | 60 |
| Dynamic range (ADC) |  | 10 bit |
| Gain (ADC) |  | min. 0 - 12 dB |
| Integration time |  | 90 µs – 20 sec |

| Radiation-hardened Camera requirements |  |  |
| --- | --- | --- |
| Type |  | CID video cameras (monochrome) for special installations, where high radiation hardness is necessary |
| Data interface |  | Video output |
| Chip resolution | pixel | 786 x 603 |
| Refresh rate | fps | 25 |
| Gain |  | 2 |

### µTCA

Table 2: MTCA.4  form factor 19” rack mount

| HW Component | Type |  | Remarks |
| --- | --- | --- | --- |
| µTCA System | GSI-NATIVE-R2-G64-AC Or RackPak/M5-1with Concurrent Technologies AM G64/471-51 CPU or better | Intel CPU 600W Power Supply 6-Slot MTCA.4 Chassis 2U NAT MCH | NAT and Powerbridge |
| Timing | AMC FTRN |  | Cosylab |
| Switch | AMC217 | 8-port Switch AMC | Vadatech |

Figure 15: Top: Proposed MTCA.4 chassis incl. MCH, 600W PS, CCT Core i7 or later CPU (NAT-Native-R2). Bottom: CPS 8 unit.

### Industry PC (IPC)

In case only few channels have to be read out (up to 2), a standard IPC can be used to host the GbE LAN adapter (e.g. dual port) and the FTRN. The IPC operates as standard FEC and will boot via PXEboot its OS.

| HW Component | Type |  | Remarks |
| --- | --- | --- | --- |
| Dual Port LAN Adapter | PCIe |  | Intel |
| Timing | Pexaria FTRN |  | GSI |
| IPC | Sysgen, Kontron |  | Core I7 or better |

### Remote Control

For maintenance purposes the MTCA system can be controlled via LAN or IPMI protocol. It provides access to status, temperature, fan speed and powercycle. The system can be surveyed by the REMBRANDT system.

The same type of remote operation is valid for the IPC.

A remote power cycle for the cameras is feasible using the Profinet based Remote Power Control (RPC) system providing a reset pulse to the CPS8 unit.

### CPS8

Power supply and remote reset for up to eight digital cameras are realized by the in-house developed Camera Power Supply controller CPS8. The CPS8 is based on an Arduino single-board microcontroller and controlled via Ethernet. By sending simple ASCII commands the controlling FESA class sets the power of any attached camera. It can also obtain the power status of each connected camera. In addition, the CPS8 distributes hardware triggers to the cameras either from an external input or self-generated.

Figure 16: Photo CPS8 system.

CPS8:

- build in-house

- Atmel Microcontroller (Arduino) based

- simple ethernet based access

- camera power control

- camera power status

- camera trigger

- software trigger

- hardware trigger distribution

- one camera = one trigger

- options like burst mode can be easily implemented

- front panel indicators

- camera power

- camera exposure (camera strobe output)

### PLC System

A Siemens PLC PROFINET IO System (IO controller and devices) controls the aperture of each lens and LEDs to illuminate the target for calibration issues.

A dedicated FESA class using the CERN SILECS library provides access to the PLC from the CUPID GUI via the ACC-Network.

The system will be set up on a 19” DIN-rail holder either close to the SCR DAQ system or in radiation-protected locations inside the tunnel. It is part of the campus PROFINET infrastructure. Plan is, to install a separate stand-alone PLC system for remote power control for all devices.

The bill of material and a schematic overview of the IO System are listed below.

| Description | Order number | Info |
| --- | --- | --- |
| SITOP PSU8200, DC 24V/5 A | 6EP3333-8SB00-0AY0 | Power supply |
| CPU 1512SP-1 PN | 6ES7512-1DK01-0AB0 | IO Controller (CPU) |
| CP 1542SP-1 | 6GK7542-6UX00-0XE0 | CP (ACC-Network) |
| BusAdapter 2xRJ45 | 6ES7193-6AR00-0AA0 | Bus Adapter |
| DQ 8x24VDC/0,5A HF | 6ES7132-6BF00-0CA0 | Digital Out - LED |
| AQ 4xU/I ST | 6ES7135-6HD00-0BA1 | Analog Out - Apertur |
| Extras |  |  |
| Miniature Circuit Breaker | 5SJ4111-7HG40 | MCB |

Figure 17: Scheme of PLC system for remote-control iris and LED control.

Figure 18: Photo of PLC system for remote-control iris and LED control.

## Supported Infrastructure (HV, Pressurized Air, Detector Gas)

The SCR detectors do require additional infrastructure. Almost all screens are mounted on movable pneumatic drives. The lens system in addition requires PLC control, so all those systems belong to the Profinet environment and must be addressed there.

# Interfaces

## Interface to Accelerator Control System

A dedicated FESA class [2,3,4,6] running on the Front-End Controller (FEC) acts as a server for all required properties. Any client such as a GUI may subscribe to these properties to GET and SET data. This client-server architecture requires a network middleware, here RDA (Remote Data Access).

The Profinet infrastructure requires the SILECS [10] gateway, a software interface between FESA and PLC systems.

## White Rabbit Timing

Any of the SCR DAQ systems require connection to the White Rabbit timing system [7]. Depending on the chosen DAQ form factor, a suitable Fair Timing Receiver Node (FTRN [8]) has to be integrated. The cameras are triggered via the CPS8 directly from the FTRN. The FTRN settings for the trigger are handled via the SCR FESA class.

## Beam Transmission Monitoring

The system will not be included into the beam transmission monitoring system.

## Machine Protection

The system will not be part of the machine protection system.

## Accelerator and Beam Modes

It is anticipated, that the system will handle the proposed Accelerator and Beam Modes. More details can be found in [9]. The implementation of these modes has to be discussed with CSCO.

## Archiving

It is anticipated that the SCR system will be included into the archiving system.

## Alarming

The SCR system will not be part of the alarming system.

## Expected Data Supply from LSA

The system can be connected to the LSA settings management [5], if necessary.

So far settings management is not applicable for the scintillating screens.

## Miscellaneous (Interlock etc.)

---

# Instrument Front-End Software

## High-Level Design

A FESA class for IPC and MTCA based camera readout is available and can be used.

## Device Drivers and Firmware

- Implementation of the CCD into the FESA class requires the vendor CCD library.

- CPS8 unit is based on individual firmware/software, SCR FESA class will communicate with the CPS8 via LAN.

- Pexaria, AMC FTRN and the Dual Port LAN adapter require compatible drivers.

## Operating System

The FEC (AMC or IPC) will load its OS via PXEboot over the network. As OS an up to date CentOS version with RT patch (kernel.org) is foreseen.

# Expected Performance

## Important Performance Parameters

Data Acquisition on the Frontend: Server for Multiple clients.

Image Acquisition and Processing

- image acquisition

- software crop

- rotation (90°, 180°, 270°)

- mirror

- simple image scaling

- projections

- simple analysis neglecting

further perspective distortion

- histogram

- position of brigthest pixel

- center-of-brightness

- FWHM

Camera / Software Settings (if supported)

- acquisition mode (free running / triggered)

- exposure time

- binning

- gain boost

- contrast

- brigthness

- rotation / mirror

- area-of-interest (fixed during runtime)

## Expert-GUI

An expert GUI is ready and operational.

Figure 19: CUPID GUI.

# Bill of Material

## Following CUPID DAQ components are required:

### Following CUPID µTCA DAQ components are required:

| Electronic room | Units | System | Crate | CPU | MCH | Power Supply | Timing | Switch |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| G004a G0704A.E20.010 | 4 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| G017.1 H0719A.E30.123 | 10 | 16 | 2 | 2 | 2 | 2 | 2 | 2 |
| G018 L0516A.E10.130 | 3 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| G007 (CR) H0209A.E10.044 | 5 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| G007 (Injection) H0209A.E10.044 | 5 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| G007 (Extraction) H0209A.E10.044 | 5 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| HESR (Süd) | 1 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| K923A.U30.160 | 1 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| K923A.U30.134 | 1 | 8 | 1 | 1 | 1 | 1 | 1 | 1 |
| Total | 35 | 80 | 10 | 10 | 10 | 10 | 10 | 10 |

**CR Order should be clarified

***HESR order should be done FZJ and it is not included in this table

****spare parts:

Shopping Card for CUPID µTCA DAQ (Status 06.06.2019)

| Name | Pcs. required | Order number | Price / pcs | Price / total |
| --- | --- | --- | --- | --- |
| CPU | 10 | CCT AM G64/471-51 |  |  |
| Crate | 10 | 6-Slot MTCA.4 |  |  |
| MCH | 10 | Nat-Phys-80 |  |  |
| Power Supply | 10 | NAT600D |  |  |
| Timing | 10 | FTRN (AMC) |  |  |
| Switch | 10 | Vadatech AMC217 |  | - € |
|  |  |  | Total | € |

### Following CUPID IPC DAQ components are required:

| Electronic room | Units | System | IPC | 2-port LAN Adapter | Timing |
| --- | --- | --- | --- | --- | --- |
| G006c K0321A.E30.016 | 1 | 1 | 1 | 1 | 1 |
| K923A.U30.160 | 1 | 8 | 1 | 1 | 2 |
| K923A.U30.134 | 1 | 8 | 1 | 1 | 2 |
| Total | 35 | 80 | 3 | 3 | 3 |

Shopping Card for CUPID µTCA DAQ (Status 06.06.2019)

| Name | Pcs. required | Order number | Price / pcs | Price / total |
| --- | --- | --- | --- | --- |
| IPC | 3 | Intel X86 (Xeon,Core I7 or better) |  |  |
| 8-channel GbE Switch | 3 |  |  |  |
| Timing | 3 | FTRN (PCIe) |  |  |
|  |  |  | Total | € |

## Following PLC Slow Controls components are required for each electronic room

| Electronic room | Units | System | Power Supply | CPU | CP | Digital Out | Analog Out | Base Unit black |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| G004a G0704A.E20.010 | 4 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| G017.1 H0719A.E30.123 | 10 | 16 | 1 | 1 | 1 | 2 | 4 | 4 |
| G018 L0516A.E10.130 | 3 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| G007 (CR) H0209A.E10.044 | 5 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| G007 (Injection) H0209A.E10.044 | 5 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| G007 (Extraction) H0209A.E10.044 | 5 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| HESR (Süd) | 1 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| K923A.U30.160 | 1 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| K923A.U30.134 | 1 | 8 | 1 | 1 | 1 | 1 | 2 | 1 |
| Total | 35 | 80 | 9 | 9 | 9 | 10 | 20 | 12 |

*for RadHard Camera (PT01DF1) no PLC slow controls will be required.

**CR Order should be clarified

*** HESR order should be done FZJ and it is not included in this table

**** spare parts included

Shopping Card for CUPID PLC Slow Controls (Status 06.06.2019)

| Name | Pcs. required | Order number | Price / pcs | Price / total |
| --- | --- | --- | --- | --- |
| Trainer Packet SCE (including 1xCPU, 2xDigital Out | 9 | 6ES7512-1SK00-4AB2 | 677,00 € | 6.093,00 € |
| Power Supply | 9 | 6EP3333-8SB00-0AY0 | 129,00 € | 1.161,00 € |
| CP (ACC- Network) | 9 | 6GK7542-6UX00-0XE0 | 470,00 € | 4.230,00 € |
| Analog Out | 20 | 6ES7135-6HD00-0BA1 | 194,00 € | 3.880,00 € |
| Base Unit-Type A0, black | 12 | 6ES7193-6BP00-0BA0 | 13,36 € | 160,32 € |
| Extras |  |  |  | - € |
| Miniature Circuit Breaker | 18 | 5SJ4111-7HG40 | 20,30 € | 365,40 € |
|  |  |  | Total | 15.889,72 € |

## Following CUPID CPS8 components are required for each electronic room

| Electronic room | Units | System | Pcs. required |
| --- | --- | --- | --- |
| G004a G0704A.E20.010 | 4 | 8 | 1 |
| G017.1 H0719A.E30.123 | 7 | 16 | 2 |
| G018 L0516A.E10.130 | 3 | 8 | 1 |
| G007 (CR) H0209A.E10.044 | 5 | 8 | 1 |
| G007 (Injection) H0209A.E10.044 | 5 | 8 | 1 |
| G007 (Extraction) H0209A.E10.044 | 5 | 8 | 1 |
| HESR (Süd) | 1 | 8 | 1 |
| Total | 30 | 80 | 8 |

*for RadHard Camera (PT01DF1, T1X1DF1, SIS18 Dump, SIS100 Dump) and SIS100 no CPS8 will be required.

**CR Order should be clarified

*** HESR order should be done FZJ and it is not included in this table

**** min. 2 spare CPS8 will be build, additionally CPS8 request for BIF/IPM should be clarified

Shopping Card for CUPID CPS 8 (Status 06.06.2019)

| Name | Pcs. required | Order number | Price / pcs | Price / total |
| --- | --- | --- | --- | --- |
| CPS8 | 8 |  |  |  |

## Following CUPID IDS digital cameras are required for each electronic room

| Electronic room | Units | System | Pcs. required |
| --- | --- | --- | --- |
| G004a G0704A.E20.010 | 4 | 4 | 8 |
| G017.1 H0719A.E30.123 | 7 | 7 | 7 |
| G018 L0516A.E10.130 | 3 | 3 | 3 |
| G007 (CR) H0209A.E10.044 | 5 | 5 | 5 |
| G007 (Injection) H0209A.E10.044 | 5 | 5 | 5 |
| G007 (Extraction) H0209A.E10.044 | 5 | 5 | 5 |
| HESR (Süd) | 1 | 5 | 5 |
| Total | 30 | 30 | 30 |

*we order already 20 IDS camera for HEBT (green), including spare parts

**CR Order should be clarified

*** HESR order should be done FZJ and it is not included in this table

Shopping Card for CUPID IDS digital cameras (Status 06.06.2019)

| Name | Pcs. required | Order number | Price / pcs | Price / total |
| --- | --- | --- | --- | --- |
| UI-5240SE-M-GL Rev.2 | 20 (HEBT) | 4500178303 | 579,50 € | 11590,00 € |
| UI-5240SE-M-GL Rev.2 | 5 (CR) |  |  |  |
| UI-5240SE-M-GL Rev.2 | 5 (pbar+CR Injecton) |  |  |  |
|  |  |  | Total |  |

## Following CUPID MRad3 Cameras and Pleora frame grabber are required for each electronic room:

| Electronic room | Units | System | Pcs. required MegaRad3 | Pcs. required Pleora |
| --- | --- | --- | --- | --- |
| G017.1 H0719A.E30.123 | 3 | 3 | 3 | 3 |
| G006c K0321A.E30.016 | 1 | 1 | 1 | 1 |
| K923A.U30.160 | 1 | 1 | 1 | 1 |
| K923A.U30.134 | 1 | 1 | 1 | 1 |
| Total | 6 | 6 | 6 | 6 |

| Electronic room | Units | System | Pcs. required MRad3 | Pcs. required Pleora |
| --- | --- | --- | --- | --- |
| G017.1 H0719A.E30.123 | 3 | 3 | 3 | 3 |
| G006c K0321A.E30.016 | 1 | 1 | 1 | 1 |
| K923A.U30.160 | 1 | 1 | 1 | 1 |
| K923A.U30.134 | 1 | 5 | 5 | 1 |
| Total | 6 | 6 | 6 | 6 |

*we order already 1 MRad3 for HEBT T1X1DF1 and 1x for pBar Target  PT01DF1 cameras (green)

** we already purchase Pleoras inkl. spare parts

Shopping Card for CUPID MRad3 Cameras (Status 25.06.2019)

| Name | Pcs. required | Order number | Price / pcs | Price / total |
| --- | --- | --- | --- | --- |
| 8726DX7 Camera MRAD3 CCIR + 200m Assembly cable and power supply | 1 (T1X1DF1) 1 SIS18 Dump 1 SIS100 Dump | 4500158869 4500 4500 | 9800,00 € 10980, 00 € 10980, 00 € | 9800,00 € 10980, 00 € 10980, 00 € |
| 8726DX7 Camera MRAD3 CCIR + 200m Assembly cable and power supply | 1 (pBar Target) | 4500158869 | 9800,00 € | 9800,00 € |
| 8726DX7 Camera MRAD3 CCIR + 200m Assembly cable and power supply | 2 SIS100 | 4500 | 10480, 00 € | 20960, 00 € |

# References

- F-DS-BD-16e “Detailed Specification for Scintillating Screens for HEBT”

- F-DS-BD-40e “Detailed Specification of the FAIR Beam Diagnostic Data Acquisition System”

- F-DG-C-01e “FAIR Control System Development Guideline “FESA Development Guideline”

- F-TG-C-02e “ACS Equipment Interfaces”

- F-DS-C-03e “Setting Generation Framework LSA”

- F-DG-C-04e “FAIR Control System Development Guideline “Equipment Integration Guideline”

- F-DS-C-05e “Timing-System”

- F-DS-C-06e ”Timing-Receivers”

- F-TC-C-07 ”Accelerator and Beam Modes”, Technical Concept, R. Bär, R. Steinhagen

- F-TN-C-13e “Interfacing of PLC Systems with FESA Based Front-End Systems”

# Open Issues

- the form factor per device is not yet decided.

- CR names mismatch with BEL Nomenlist

# List of Abbreviations

ACS			Accelerator control system

ADC			Analog to Digital Converter

AMC			Advanced Mezzanine Card

AMT			Active Management System

API			application programming interface

BD/BI			Beam Diagnostics, Beam Instrumentation

BLM			Beam Loss Monitor

COTS			Commercially off-the-shelf

CR			Collector Ring

Cu			Copper

DAQ			Data acquisition

DCT/DCCT		DC Current Transformer

nDCCT		Novel DCCT

FAIR			Facility for antiproton and ion research

FCT			Fast Current Transformer

FEC			Front-end Computer

FESA			Front-end software architecture

FMC			FPGA mezzanine card

FPGA			Field Programmable Gate Array

FTRN 			FAIR Timing Receiver Node

GbE			Gigabit Ethernet

GSI			Helmholtz Centre for Heavy Ion Research

GUI			Graphical user interface

HEBT			High energy beam transfer lines (at FAIR)

HESR			High Energy Storage Ring

HEST			Hochenergiestrahlführung (existing GSI)

HPC			High Pin Count (FMC Bus)

HV			High Voltage

I/O			Input Output (Register)

IPC			Industrial Personal Computer, 19" Rack mount housing

IPMI			Intelligent Platform Management Interface

IT			Information Technology

JAPC			Java API for parameter control

LPC			Low Pin Count (FMC Bus)

LSA			LHC software architecture

MTCA/µTCA		Micro Telecommunications Computing Architecture

NIM Bin		Crate for pulse shaping electronics

NIM			Nuclear Instrumentation Standard, Pulse level

PCI			Peripheral Component Interconnect

PCIe			Peripheral Component Interconnect Express

PICMG		PCI Industrial Computer Manufacturers Group

PLC			Programmable Logic Controller

PMC			PCI mezzanine card

Profinet		Process Field Network

PXE			Preboot eXecution Environment

REMBRANDT		Remote Beam Instrumentation And Network Diagnosis Tool

RPC			Remote Power Control

RT			Resonant Transformer

SIS18/100		Schwerionensynchrotron

SW			SoftWare

TMR			Tunnel Magneto Resistance

TTL			Transistor Transistor Level, Pulse level

UHV			Ultra High Vaccum

UPS			Uninterruptable Power Supply

VME			VERSA module Eurocard